# Phraxis InterGraft System

Contents

1	The InterGraft System	0

2	Pre-Procedure Checklist	0

2.1	Required Supplies	0

3	InterGraft System Procedure	0

3.1	Prepare Arterial InterGraft	0

3.1.1	Deliver Arterial InterGraft	0

3.2	Prepare Venous InterGraft	0

3.2.1	Deliver Venous InterGraft	0

# 1	The InterGraft System

The InterGraft system is a minimally invasive procedure for delivering an arteriovenus graft.

The Phraxis InterGraft System connects a standard 6-millimeter arteriovenous graft without sutures. The connectors are delivered via a catheter, and they employ stent-like technology to provide both venous and arterial flow to the graft. Unlike a standard sutured graft, the InterGraft System is designed to provide a smooth, controlled flow.The InterGraft System is a minimally invasive procedure that takes about 30 minutes, and can provide next-day vascular access for hemodialysis.

# 2	Pre-Procedure Checklist

## 2.1	Required Supplies

Prepare these items prior to the InterGraft procedure

Example

- Two micro puncture kits (needle, access wire, 4F sheath)

- One 11F introducer sheath (maximum length 11 cm)

- One 7F introducer sheath (maximum length 11 cm)

- One 0.018" diamter soft tip guidewire (length 145 cm)

- One 0.014" diameter soft tip guidewire (length 145-185 cm)

- Three stopcocks

- Graft tunneling set

- Atraumatic vascular graft clamp

- One each sterile 3 cc and 10 cc syringe

- Sterile saline solution

- Sterile 50:50 heparinized saline for flush

- Contrast agent

- 6 mm sterile stright graft (non-tapered), length appropriate for planned AV access

# 3	InterGraft System Procedure

## 3.1	Prepare Arterial InterGraft

Arterial InterGraft system preparation

Follow these steps prior to delivering arterial InterGraft connection

- Using sterile technique, open AIG package and place InterGraft delivery system in a sterile field.

- Remove the white graft grippers.

- Attach stopcocks to each of the two T-B ports A, F and open the valves B, C.

  - Close valve C and flush sterile saline through port F.

  - Close stopcock on port F and remove 10 cc syringe.

  - Close valve B and flush sterile saline through port A.

  - Close stopcock on port A and remove 10 cc syringe.

  - Flush guidewire lumen through port G using sterile saline.

  - Gently tighten T-B valves B, C.

- Attach stopcock to balloon luer E. Evacuate balloon using a 10 cc syringe, close stopcock and remove syringe.

  - Draw up 50:50 contrast/saline soultion in a 3 cc syringe, attach to stopcock, open stopcock and inject solution.

  - Visually confirm positioning balloon inflation.

  - Withdraw contrast from the balloon, close the stopcock, and leave 3 cc syringe connected to luer.

- Loosen T-B valves B, C, backload 0.014" wire into AIG delivery system and gently tighten T-B valve B.

- Deliver Arterial InterGraftArterial InterGraft implant procedure

### 3.1.1	Deliver Arterial InterGraft

Arterial InterGraft implant procedure

Follow these steps to implant the aterial InterGraft connection

- Prepare Arterial InterGraftArterial InterGraft system preparation

- Identify arteriotomy site and mark with surgical clamp.

- Attach vessel loops or vascular clamps for proximal and distal control of artery. Do not tighten at this stage.

- Gain access to artery using a micropuncture kit, insert wire, remove needle, and insert 4F sheath.

- Exchange 4F sheath for 7F sheath - replace micropuncture wire with 0.0014" wire.

- Load AIG delivery system over the 0.014" wire.

- Under fluoroscopic guidance, advance AIG over the wire and into the artery approximately 1 - 2 cm past arteriotomy.

Arrow points to clamp tip at arteriotomy site.

- Remove red deployment lock labeled D and position AIG marker band approximately 2 cm central to clamp tip that marks the arteriotomy site.

- Loosen T-B valve B.

- Reconfirm marker band position, remove 7F sheath and deploy AIG under floroscopy using standard "pin and pull" technique until only the tines emerge from the delivery sheath.

- Inflate positioning baloon.

  - Pull the inflated baloon back to engage tines, tighten T-B valve B.

  - Slowly pull entire AIG system up vertically from artery until resistance is felt.

Resistance indicates apposition of the tines at the artery wall.

- Complete deployment of the AIG using "pin and pull" technique.

  - Deflate balloon.

  - Remove entire system and guidewire.

- Grasp protruding end of AIG to control bleeding.

  - Stabilize base of AIG by grasping.

  - Insert AIG into graft.

  - Remove clamp and verify flow through circuit.

- Release vessel loops or vascular clamps, close incisions.

- Prepare Venous InterGraftVenous InterGraft delivery system preparation

## 3.2	Prepare Venous InterGraft

Venous InterGraft delivery system preparation

Follow these steps prior to delivering venous InterGraft connection

- Using sterile technique, open VIG package and place InterGraft delivery system in sterile field.

- Remove the white graft grippers.

- Attach a stopcock to each of the T-B ports D, E.

- Loosen the T-B valves A, C.

- Using 10cc syringe, flush the delivery system with sterile saline solution. Close valve A

- loosen both T-B valves A, C.

  - Load the 0.018" guide wire into the VIG delivery system.

  - Backload wire into delivery system and gently close T-B valve C.

- Deliver Venous InterGraftVenous InterGraft implant procedure

### 3.2.1	Deliver Venous InterGraft

Venous InterGraft implant procedure

Follow these steps to implant the venous InterGraft connection

- Prepare Venous InterGraftVenous InterGraft delivery system preparation

- Pre-tunnel graft with ends adjacent to target vein and artery anastomotic sites using specified tunneling procedure.

- Identify target venotomy site and mark with a surgical clamp.

- Access the vein using micropuncture kit, insert wire, remove needle and insert 4F sheath.

- Exchange 4F sheath for 11F sheath--replace micropuncture wire with 0.018" wire.

- Load VIG delivery system over the wire.

- Under flouroscopic guidance, advance VIG over the wire and into vein, approximately 1 cm past venomy.

- Remove blue deployment lock labeled B and position marker band approximately 1cm peripheral to clamp.

Arrow points to marker band.

- Loosen T-B valve C and under fluoroscopic guidance slowly advance the hypotube until it just engages the VIG or

"closes the gap"

- Reconfirm the marker band position and remove the 11F sheath from vein.

- Deploy VIG under floroscopy using standard "pin and pull" technique, ensuring the marker bands and venotomy clamp remain in position throughout deployment.

- Remove VIG delivery system and wire.

  - Gently grasp and compress protruding end of VIG to control bleeding.

  - Insert VIG into pre-tunneled graft end.

- Verify blood flow through graft

  - Clamp graft to stop venous back-bleeding

  - Prepare 10 cc syringe with hepranized flush

  - Remove graft clamp

  - Flush graft

  - Reattach clamp